SkyGuard-X

AWS Anomaly Intelligence — Simple User Manual

Smart India Hackathon 2026 · Problem Statement SIH26073 (MoES / IMD) · Prototype v2 · 13 Sep 2026

SkyGuard-X watches Automatic Weather Station (AWS) data — temperature, pressure and humidity — and tells you, for every reading, whether it is normal, a sensor fault, a communication problem or a genuine weather event. For each alert it gives the cause, severity, confidence, a plain-English explanation, a corrected value and the maintenance action to take. This manual explains how to install it, use the website, upload your own data and read the results.

# 1. What you need

A laptop with Windows 10/11, macOS or Linux (4 GB RAM is enough).

Python 3.10 – 3.13 from python.org. On Windows tick “Add Python to PATH” during installation.

The zip file skyguard_x_prototype_v2_2026-09-13.zip (about 28 MB).

A web browser (Chrome, Edge or Firefox). No internet is needed after installation.

# 2. Install and start (5 minutes)

Unzip the file. You get a folder named skyguard.

Open a terminal inside that folder (Windows: type cmd in the folder’s address bar and press Enter; macOS: right-click the folder → “New Terminal at Folder”).

Run the commands below for your system. The first install takes 2–5 minutes.

When you see “Uvicorn running on http://0.0.0.0:8000”, open http://localhost:8000 in your browser.

Windows (Command Prompt / PowerShell):

python -m venv .venv
.venv\Scripts\activate
pip install -r requirements.txt
python run.py

macOS / Linux:

python3 -m venv .venv
source .venv/bin/activate
pip install -r requirements.txt
python run.py

To stop the server press Ctrl + C in the terminal. Next time you only need to activate the environment and run python run.py again.

# 3. The website — five tabs

The top-right buttons switch between the five pages. Everything runs on your laptop; nothing is uploaded to the internet.

## 3.1 Network — the big picture

Network tab: KPIs, health map, latest alerts and the station table.

Four KPI cards: readings scored, sensor faults found, communication/data issues, and genuine weather events (readings that looked strange but neighbours confirmed were real).

Map: 43 real IMD/WMO stations. Colour = health (green healthy ≥ 80, yellow watch, orange maintenance, red critical). Dot size = number of faults. Lines join each station to the neighbours used for the spatial check. Click a dot to open that station.

Latest alerts: the most recent problems across the whole network, with cause, severity and confidence.

Stations table: health 0–100 for the station and for each sensor (T, P, RH), counts of faults, comms issues and genuine weather events, and the time of the last alert.

## 3.2 Station — replay one station

Station tab: observed vs expected values with flagged readings marked; click an alert row to see why.

Choose a station from the drop-down and a window (7 / 14 / 30 / 60 days).

Charts: solid line = what the sensor reported, dashed line = what SkyGuard-X expected (from climatology and neighbours), red dots = readings flagged as faulty.

Health cards: current health of temperature, pressure and humidity sensors with status (HEALTHY / WATCH / MAINTENANCE / CRITICAL).

Alerts in window: click any row to expand the plain-English explanation, the recommended action, the corrected value, the bias estimate and the top reasons (SHAP contributions) behind the decision.

## 3.3 Judge mode — upload your own CSV (the main demo)

Judge mode: a messy 15-minute file with °F temperatures and semicolons — spike, frozen humidity and a 5-hour gap all found, zero false alarms.

This is where an evaluator tests the system on their own anomaly-injected data.

Click “Judge mode · upload CSV”.

Drag a file into the dashed box, or click browse. Accepted: .csv (comma, semicolon, tab or pipe separated), .tsv and .xlsx.

Wait a few seconds. The status line shows rows scored and speed (about 1–3 ms per row).

Read the Run report (right): verdict counts, how your columns were mapped, automatic fixes (units, timestamp format), detected cadence, unseen stations, faults by cause and per-station health.

Scroll down: the chart shows your uploaded station (switch stations with the drop-down) and the Alerts table lists every flagged reading. Click a row for the explanation.

Click “Download scored CSV” to get your file back with the verdict columns added (see section 5).

What your file needs:

A timestamp column (any common format: 2025-04-03 02:00, 03/04/2025 02:00, ISO, or Unix epoch).

Temperature, pressure and humidity columns — any names (temp / Temperature / T / air_temp / TEMP_F …) and any units (°C, K, °F; hPa, Pa, inHg; %). They are recognised and converted automatically.

A station column is optional. Without it the file is treated as one station.

Fill values (-999, 9999 …), duplicate timestamps, gaps and any cadence (1-minute to hourly) are handled.

Don’t have a file? Click “Download a messy sample CSV” to get a test file with Kelvin, Pascals, day-first dates and sentinels, then upload it.

Tip: unknown stations work too — SkyGuard-X fits a climatology from the file itself (“cold start”), so it can be used on any AWS network, not just the 43 built-in stations.

## 3.4 Benchmarks — the numbers

Benchmarks tab: honest, held-out results on 169,248 test hours with injected faults.

Event detection rate, precision / recall, false-positive rate on normal hours and false-alarm rate on genuine extreme weather.

A table per fault type (spike, frozen, drift, bias, noise, inconsistency, missing, fill, stale) with detection rate and time-to-detect.

All numbers come from a chronological hold-out (2025) the model never saw during training — quote them as they are.

## 3.5 How it works

A one-page summary of the pipeline (Tier-1 deterministic QC → 98-feature context engine → LightGBM detector and root-cause → verdict layer → explanations) and what makes it different: real data, the Hubbard Spatial Regression Test, witness agreement for genuine weather, physics features, CUSUM drift tracking, cold start and the ESP32 edge tier.

# 4. Reading a result

Verdicts

| Verdict | Meaning | What to do |
| --- | --- | --- |
| SENSOR_FAULT | The instrument is wrong (spike, frozen, drift, bias, noise, T/RH inconsistency, out-of-range). | Follow the action shown; use the corrected value. |
| DATA_COMM_ISSUE | Telemetry problem: missing data, fill/sentinel values, stale repeated packets, duplicate timestamps. | Check power, modem/link, logger clock or firmware. |
| GENUINE_WEATHER_EVENT | Reading is unusual but neighbouring stations agree — real weather (thunderstorm, heat wave, cyclone). | No maintenance. Do not reject the data. |
| UNCERTAIN | Some evidence of a fault, below the alert threshold. | Watch; usually resolves to NORMAL or becomes a fault as it persists. |
| NORMAL | Nothing wrong. | — |

Cause (root cause) and action

| Cause | Typical sign | Recommended action shown by the system |
| --- | --- | --- |
| spike | One-point jump that reverts | Transient – auto-reject reading, keep monitoring |
| frozen | Identical value for hours while others move | Sensor stuck – inspect element/connector; likely replacement |
| drift | Slow growing offset vs neighbours/climatology | Schedule recalibration |
| bias | Sudden constant offset | Verify against travelling standard; recalibrate |
| noise | Jittery readings | Check cable/shield & logger ADC |
| inconsistency | RH/T relationship broken (dew point > T) | Check RH element (wetting/contamination) |
| missing / fill / stale | Gaps, -999/9999 values, repeated packets | Check power, link, logger firmware or clock |

Severity and confidence

Severity LOW → MEDIUM → HIGH → CRITICAL, based on how far the reading is from what is expected (roughly >3σ, >5σ, >8σ).

Confidence 0–1: the model’s probability that the reading is faulty (alerts are raised at ≥ 0.85; Tier-1 rule hits are 0.95–1.0).

Corrected value the best estimate of the true reading, blended from neighbouring stations and the station’s own climatology.

Bias estimate the current offset of the sensor (e.g. +3.1 °C), useful for recalibration.

Health 0–100 per station and per sensor over the last 7 days: ≥80 HEALTHY, 60–79 WATCH, 40–59 MAINTENANCE, <40 CRITICAL.

# 5. Command line (batch scoring, no browser)

python -m skyguard --input yourfile.csv --output scored.csv
python -m skyguard --input yourfile.xlsx --output scored.csv --no-explain   # faster, no SHAP text

The command prints a JSON run report and writes scored.csv with your original columns plus:

station_id, ts, temp, pres, rh, verdict, fault, sensor, severity, confidence, p_fault,
corrected, bias_estimate, temp_expected, pres_expected, rh_expected, explanation, action

Docker (optional, e.g. for Hugging Face Spaces):

docker build -t skyguard .
docker run -p 7860:7860 skyguard      # then open http://localhost:7860

# 6. Suggested 3-minute demo

Network tab – point at the map: “43 real IMD stations, colour is health, lines are the neighbour links used for the spatial test.”

Click a red dot → Station tab – show the observed vs expected lines and click one alert to reveal the explanation, action and corrected value.

Judge mode – upload the evaluator’s file (or the messy sample). Show that units and columns were auto-mapped, then the verdict chips and the alert list. Click the spike row: “58 °C, 87σ from climatology, no neighbour supports it → sensor fault, corrected to 23 °C”.

Benchmarks tab – read the real numbers: 86 % of injected fault events detected, 0.49 % false positives on normal hours, 0.69 % false alarms on genuine extreme weather.

Download the scored CSV and hand it over.

# 7. Troubleshooting

| Problem | Fix |
| --- | --- |
| “python is not recognized” (Windows) | Use py instead of python, or reinstall Python with “Add to PATH” ticked. |
| pip install fails on lightgbm / shap | Use Python 3.11 or 3.12 (newest Python versions may not have wheels yet). |
| macOS: libomp.dylib not found | Run brew install libomp, then start again. |
| Port 8000 already in use | python -m uvicorn app.main:app --port 8080 and open http://localhost:8080 |
| First page takes ~4 minutes to load | data/demo_scored.parquet was deleted; it is being rebuilt. Keep the file. |
| Upload says “Could not identify columns” | Make sure the file has a timestamp and three value columns; rename headers to temp / pres / rh if they are cryptic. |
| Upload says “file too large” | The demo server accepts up to 400,000 rows per upload; split larger files or use the command line. |
| Charts empty after upload | Timestamps could not be parsed — use a standard format such as 2025-04-03 02:00. |

# 8. Known limits (be upfront)

Small single-station offsets (below about 4 °C) are only caught after they persist for several hours; noise is the weakest class (≈59 % of events).

Spatial checks need at least 2–3 neighbours within reach; a lone station relies on temporal, climatology and physics tests only.

This is a hackathon prototype on public archive data, not an operational IMD product.